## Slide 1: Filter Task Force
Filter improvement:
    Oversampled filter bank
Stakeholders :
    Radio Frequency environment
    Calibration: (Spectral line, Continuum source)
    Pulsar
Test Bench:
    Simulation: Matlab, modelsim, Fpga

## Slide 2: Polyphase filter structure used
To simulate the polyphase filter bank : 32MHz effective bandwidth
Simulate the filter bank sweep: from 10MHZ to about 300 MHz
1 polyphase filter bank is composed by a N subbands of FIR filter corresponding to the modulation of the prototype filter

## Slide 3: Reinterpolation and filter
First obtain the frequency response of the filter by FFT and interpolation to get the channel resolution
Fit a rectangular window on the filter bandpass in order to study the behaviour of the filtered data
Prototype filter after modulation : channel 40

## Slide 4: Filter measurement
Reference : “Downsampling simulation” with a perfect rectangular window in making the linear mean of all the channels
Modulation of the filter across the band. Multiplication with data point in the frequency domain and summation of the power obtain and divided by the same number of channel as the reference

## Slide 5: Result
Use of the filter
Reference

## Slide 6: Result(2)
In making computing the quadratic error (linear scale) between the 2 plots we obtain the final result, the residual power left by the filtering
These results are obtained for 1 polyphase filter bank

## Slide 7: Final result(1) for 8 polyphases
Reference data
FIR sweeping

## Slide 8: Final result(2)
Residual power
Input power

## Slide 9: RFI investigation
Simulations using real data:
  3 stages of filtering had the contribution: Anti-Aliasing filter before quantization. Modelize the analog filter used ?
  Stage 2: We are able after to evaluate the impact of the RFI at 256Khz channel over the full frequency range 32MHz. Choice of the window (side lobes level + big transition band impact)
  Stage 3 :Investigate the 1Khz level will tell about the RFI for a particular frequency range of the 256Khz. Risks are higher for this resolution the RFI can be from 3 to 100kHz spectrum so more chance to have the total bandwidth of 256Khz occupied, so more folding power.
Theta framework:
  FIR filter available, experiments possible. Sine wave increasing in 1 frequency study when the power is powerful enough to disturb the bandwidth of interest.

## Slide 10: FIR filter plugin
Parametrable FIR Filter plugin available in the Theta ITS framework.
Frequency
Gain in dB
Frequency
Gain in dB

## Slide 11: Calibration
Correct for the filter behavior at subband level:
    Gain : Deterministic and stable in time
      Condition: smoothness of the gain versus frequency. ( Less than 10 parameters needed to describe the gain behavior.
    Phase : Deterministic and stable in time
      Condition: Phase linearity avoiding phase jumps, (smoothness) assured by linear filter.
        Get the phase characteristic of the polyphase filter in 2 cases. Critically sampled and oversampled case
    Advantage of digital processing : filter response stable in time Evaluation of the correction factors (Gain, Phase) can be done offline and apply to the data.
    Who is doing the processing: Stations or the Core?
    Investigation to get the phase response of the filter bank

## Slide 12: Calibration
Gain => Beampattern effects : beam squint to correct as well.
Correct for the signal degradation introduce by filtering: Aliasing due to the critically sampled version.
=> The correlation can take of the problem by having a shifted version of the channels for each station.
Spectral Image
Critical sampling implies folding band
Stop band : -75db
Disappear after correlation

## Slide 13: Pulsar Experiments
  Is again the all chain has to be consider
    1 – Impact of the beamforming: Cascad of the polyphase
    2 – Impact of the filtering
  1 - Two stages of polyphase filter and 2 (1 ? Offline) stage inverse
    Beam pointed straight
    Beam pointed at a source
Polyphase
Filter 1
Polyphase
Filter 1
Polyphase
Filter 1
+
Polyphase
Filter 2
256khz*256
256khz*256
256khz*256
256*256khz
1khz*256*256
64MHz
64MHz
64MHz
Inverse polpyphase
Filter cascad
64MHz
- Variance between input and output

## Slide 14: Pulsar
Pulsar parameters :
  Dm: Dispersion measure: Phase dispersion occurs in the interstellar medium
  Rho: coherency matrix describing the polarization
  SNR: Signal to noise ratio
  P: pulse period
  Number of component: Nc gaussian intensity peaks
    For each component: height, width, phi: polarization
Idea: Characterize signal distortion as a function of polyphase filter parameters : Nsubband, Rp, Rs…
Proposal: active collaboration with pulsar specialists
  Phase 1: using a large pulse period P and high width W, features are then adequately resolved. How the data are affected by Polyphase Filter in using pulsa generator (W.van Straten, Ramesh K)
  Phase 2: We shorten P and w and test that inverting the analysis filter bank recovers the pulsar parameters
  Phase 3: extend the experiment to other parameters, polarization…..

## Slide 15: Conclusion
RFI: is the critical aspect in order to investigate the best window for the processing.
  Side lobe level shape
  transition band
Calibration: give us constraint
  on the pass band ripple
  also the overlapping caracteristic of the filter: Gain and phase
  Another effect will have to be coupled: beamsquint
Pulsar : give us constraint on
  Temporal depth which will be dependant on the oversampling factor
  Signal degradation involve in the near-perfect reconstruction filter
Issues: Not only filtering but also beamforming characteristic can not be decoupled of the study.

## Slide 16: Filter test bench
ADC : 80Ms/s
Control Pc
GPIB interface
Ethernet link 100Mb/s
Data Wrapper
Data Unwrapper
FPGA
ADC
Data generator

## Slide 17: Test bench(2)
Signal generator Marconi (GPIB interface)
A/D converter: 12 bits, 40Mhz
FPGA : VHDL Program
Signal generator Marconi (GPIB interface)
Control Pc
  Pentium 233 Mhz, 128Mo Ram / Pentium 600 Mhz (soon) because lake of computational power
  Carte GPIB : nationale Instrument
  Carte ethernet : 100 Mbits

## Slide 18: Test bench (3)
OS : Linux, Suse 8.2, Specific Kernel Compilation
Driver National Instruments
C++ Application :
  GPIB Communication
  Ethernet Client (ethernet frame: 1046 bytes): Available
  Data wrapper: (Header + data extraction) Available(need optimisation
  File recording : Available
  Find automatically the maiximum dynamic range
  Sine wave fitting: Few points to evaluate the sinewave, Gain (max value determination, Phase determination (using the input sine wave as a reference -> Matlab program available, C++ under developpement
  Status: GPIB connection, Sinwaves comes out,
  Now: fitting: Gain plot, Phase Plot (Reference signal in FPGA)

## Slide 19: Sinewave fitting technique
2 methods had been investigated
Minimum variance:
  Method based on iteration in minimizing step by step the variance between the fitting function and the sine wave. The stop criterium is based on an epsilon.
    Heavy to process
    Stop criterion not sure to reach in case of noise
Fourier domain analysis:
  Suppose vec(y) is the vector containing the data points and vec(b) is the base vector to project on with b_i = exp(j*omega*t_i).
  proj = vec(y) * vec(b)where * denotes the inner product, we find the amplitude and phase information:
  A = abs(proj) / (N/2) , phi = angle(proj) - pi/2
  Advantage very fast
  Accuracy depending on the number of data points used, larger will be the result better will be the fitting accuracy

## Slide 20: Result of the fitting method
Gain
Phase

## Slide 21: Verification of the implementation
FIR filter ModelSim VHDL simulator and Matlab:
We had differences in the implementation between Matlab and ModelSim.
The aim was to explain the differences.

## Slide 22: Quantization error propagation
Quantization of the coefficients
Quantization of the input signal
Quantization of the output signal
Error propagation due to the difference in the quantization of the input.

## Slide 23: Hardware verification
Possibility to use Matlab test vector to verify the FPGA processing for certain conditions.
  Input and output should really be treated by the same way
Towards the Polyphase Filter Implementation
  Polyphase filter Structure (several FIR Filter + FFT)
Truncation by Matlab vs ModelSim?? (Polyphase filter structure implementation + comparison)
FFT:
  Floating point. (no Matlab function for the Altera FFT). Estimate an Epsilon
  Fixed Point (Matlab function)
Aim is to give exact values or boundaries to validate the FPGA implementation.
D/A conversion problem or intermodulation product.
  To decouple the problem a sinewave generator had been integrated in the FPGA. (limitation to certain frequencies)
  Comparison possible at the same frequencies after A/D

## Slide 24: Conclusion
Stakeholders are driving the requirements: best design requirements can be found out experience.
We can already give some safe requirements. Rp, Rs
This presentation has been mainly focus on a experimental point of view:
  The experiments for each stakeholders to test our filter solution
  Experiments in order to evaluate the hardware implementation accuracy compare to the simulation: Give boundary of the reliability of the result
The hardware implementation limitations will drive the signal quality generalisation of method ?
Issues:
  We have to integrate the spatial filtering into the study due to calibration and pulsar processing. But real experiments can be carried out by the theta, ITS processing chain (calibration….)
  Documentation must takes place in january
  Oversmpled polyphase filter only implemented under Matlab and needs improvements
  Synthesis filter needs to be implemented: Matlab, C++, FPGA ??